# Inhibition of OATP1C1-mediated cellular uptake of radiolabeled thyroxine

Standard Operating Procedure

# Objective and scope

This Standard Operating Procedure (SOP) describes the application of a cell based in vitro assay to screen chemicals for their capacity to inhibit thyroxin (T4) uptake by human organic anion transporter polypeptide 1C1 (OATP1C1) using an overexpressing cell model.

# Definition

The organic anion transporter polypeptide 1C1 (OATP1C1, expressed by the SLCO1C1 gene) is a transmembrane protein that is important for cellular uptake of thyroxin (T4) in the brain. It facilitates the transport of T4 across the cell membrane in astrocytes, endothelial cells present in the blood-brain-barrier (BBB) and choroid plexus cells present in the blood-cerebrospinal fluid-barrier (BCSFB). T4 is the precursor form of the active triiodothyronine (T3) thyroid hormone that is important for the fetal neurodevelopment. Disruption of T4 transport into cells might cause TH deficiency in the brain, which can lead to severe adverse neurodevelopmental effects. Certain chemicals can interfere with the transport of T4 into cells by inhibition of OATP1C1. The disruption of the OATP1C1 facilitated thyroid hormone transmembrane transport is regarded as a possible mechanism for thyroid hormone system disruption by chemicals.

# Principle

T4 has been labelled with a radioactive isotope on the iodine group producing a radiolabeled T4 (125I-T4). This radiolabeled isotope is used for the quantification of the amount of T4. Chinese hamster ovary (CHO-K1) cells have been stably transfected with a plasmid encoding the gene for OATP1C1 and, as a result, overexpress OATP1C1 on their membrane. After exposure to 10 nM of 125I-T4 for 30 minutes, cells are washed to remove extracellular 125I-T4, and intracellular 125I-T4 is quantified by measuring the radioactivity of the cell lysate in a gamma counter. To determine the OATP1C1 inhibiting capacity of a test item, cells are co-exposed to 10 nM of 125I-T4 for 30 minutes in combination with a concentration of test item or a solvent control. If the test item is capable of inhibiting OATP1C1 mediated uptake of 125I-T4, intracellular radioactivity will be lower in cells exposed to the test item compared to cells exposed to the solvent control. If a test item causes a concentration-dependent decrease in cellular uptake of 125I-T4, a concentration-response curve is fixed to the intracellular radioactivity measurements, and its OATP1C1 inhibiting capacity is expressed as a benchmark concentration (BMC20/IC20). To exclude that a decrease in cellular uptake of 125I-T4 is due to cytotoxicity of the test item to the cells, cell viability assays are run in parallel with the same cells exposed to the test item under similar exposure conditions as in the 125I-T4 uptake experiments. Development of the assay is described by Wagenaars et al. 2024.

# Safety

The solvents and chemicals used can be toxic, harmful and/or flammable. Verify the safety codes affixed to the labels. Work in the flow cabinet, use a lab coat and gloves and refer to the Material Safety Datasheet forms.

Precautionary measures should be taken to work with radioactive chemicals and with genetically modified cells. Test system, reference item, chemicals and reagents

# Test system, reference item, chemicals and reagents

## Test system, reference item, and labels

### The test system contains of Chinese hamster ovary (CHO-K1) cells stably transfected with a pCRII-TOPO vector containing a 3.1 kb fragment encoding for human OATP1C1. These CHO-K1 OATP1C1 cells were constructed at the University Hospital of Zürich (Switzerland) and have been described by Pizzagalli et al. (2002).

### Indocyanine green (Cayman chemical, art. no. 27462 is used as reference OATP1C1 inhibitor.

### DMSO (Acros, art.no. 10438920) as reference cytotoxic compound.

### Thyroxin (Sigma Aldrich, art.no. T2376) is used as OATP1C1 substrate.

### 125I-Thyroxin (PerkinElmer, art.no. NEX111X100UC) is used as label for OATP1C1 substrate.

## Chemicals

### DMSO (Acros, art.no. 10438920)

### Ham’s F12 nutrient mix (L-glutamine) (Gibco, art.no. 11765054)

### Fetal Calf Serum (FCS) (Gibco)

### Penicillin Streptomycin solution (100x) (Gibco, art.no. 15140-122)

### Trypsin-EDTA (0.5%), no phenol red (Gibco, art.no. 15400)

### Sodium butyrate (Merck, art.no. B5887)

### Sodium chloride (Merck, art.no. S9625)

### Sodium hydroxide (Merck, art.no. S8045)

### HEPES (Merck, art.no. H4034)

### Potassium chloride (Merck, art.no. P9333)

### Calcium chloride (Merck, art.no. C1016)

### Magnesium chloride hexahydrate (Merck, art.no. M2670)

### D-(+)-glucose (Merck, art.no. G8270)

### Phosphate buffered saline (PBS) tablets (Fisher Emergo, 11455015)

### Bovine serum albumin (BSA) (Merck, B4287)

### G418 disulphate solution 50 mg/mL (AppliChem, A6798)

### MilliQ water

## Equipment and facilities

### Gamma counter (for example Wizard2, Perkin Elmer)Luminescence platereader (BioTek Cytation 5)

### Luminescence platereader (for example BioTek Cytation 5)

### Flow cabinet

### Autoclave

### Water bath at 37C

### CO2 incubator at 37C and 5% CO2

### Cell counter

### pH meter

### Pipet sets

### Vortex

### Plate shaker

### 75 cm2 culture flask (Greiner Bio-one, art.no. 658175)

### Sterile transparent tissue culture treated 96 wells plates (Greiner Bio-one, art.no. 65180)

### Sterile white tissue culture treated 96 wells plate (Greiner Bio-one, art.no. 655098)

### Syringe 50 mL (Terumo SS+50L1)

### Sterile cellulose filters 0.2 µM (Sarstedt Ref. 83.1826.102)

### Eppendorf tubes

### 10 ml serological pipettes

### 50 ml Falcon tubes

### 1.5 mL screw cap tubes (Sarstedt order no. 72.687)

### 3.5 mL polypropylene tubes (Sarstedt order no. 55.535)

### Push cap for polypropylene tubes (Sarstedt order no. 65.809)

### 1 liter glass bottles

### 500 ml glass bottles

## Reagents

### F12 Complete culture media

- Work under sterile conditions in a flow cabinet

- Open a 500 mL bottle of F-12 (5.2.2).

- Remove 50 mL from the bottle

- Add 50 ml of fetal calf serum (5.2.3) (f.c. 10%)

- Add 5.5 ml Penicillin Streptomycin solution (5.2.4)

- Mix the bottle by swirling.

- Store at 4C (stable for 1 month)

### Serum-free sodium butyrate solution (50mM) in F12 media (sterile)

- Work under sterile conditions in a flow cabinet

- Take a 50 mL falcon tube (5.3.19)

- Weigh 275.22 milligrams of sodium butyrate (5.2.6)

- Add 50 mL of serum free F12 medium (5.2.2)

- Dissolve by vortex (5.3.10)

- Using syringe (5.3.15) with sterile filter (5.3.16) filter sterilize solution and aliquot in portions of 1 mL in 1.5 mL screw cap tubes (5.3.20)

- Store at -20C (stable for 1 year)

- Dilute 1:10 in serum free F12 medium (5.2.2) freshly on day 3 of the experiment

### 0.1M NaOH in Milli-Q water

- Weigh 4 grams of NaOH (5.2.8) in a 1 Liter bottle (5.3.21)

- Add 1 Liter of MilliQ water (5.2.16)

- Mix to dissolve the pallets

- Store at room temperature (stable for 1 year)

### 1 mM Thyroxine (T4) stock solution dissolved in 0.1M NaOH

- Weigh 0.388 gram thyroxine (5.1.4) in a 50 ml falcon tube (5.3.19)

- Dissolve in 50 ml of 0.1M NaOH (5.4.3)

- Aliquot in portions of 1 mL in Eppendorf tubes (5.3.15)

- Store at -20C (stable for 1 year)

### Phosphate Buffered Saline (1xPBS)

- Add 5 PBS tablets to a 500 ml bottle (5.2.14)

- Add 500 ml Milli-Q (5.2.17)

- Mix to dissolve the tablets

- For cell culture purpose: autoclave the PBS for 20 minutes at 121C

- Store at room temperature (stable for 1 year)

### PBS + 0.1% BSA

- Work under sterile conditions in a flow cabinet

- Add 500 mg BSA (5.2.15) in a 500 ml bottle (5.3.24)

- Add 500 ml sterile 1xPBS (5.4.5)

- Store at 4C, should be ice-cold for the assay (stable for 1 month)

### Trypsin working solution

- Work under sterile conditions in a flow cabinet

    - Take a 500 ml glass bottle (5.3.24)

    - Add 50 ml trypsin stock solution (5.2.5)

    - Add 450 ml sterile 1xPBS (5.4.5)

    - Mix by shaking

    - Aliquot in 50 ml Falcon tubes (5.3.19)

    - Store at -20C (stable for 1 year)

### 1 M D-Glucose solution

- Work under sterile conditions in a flow cabinet

- Weigh 9 grams D-Glucose (5.2.13) in 50 ml falcon tube (5.3.19)

- Fill up the tube with Milli-Q water (5.2.17) till the 50 ml mark.

- Using syringe (5.3.15) with sterile filter (5.3.16), filter sterilize and aliquot in portions of 1 mL in Eppendorf tubes (5.3.15)

- Store at -20C (stable for 1 year)

### Uptake buffer: 125 mM NaCl, 10 mM HEPES, 5 mM KCl, 1.3 mM CaCl2, 1.2 mM MgCl2, pH adjusted to 7.4

- Take a 1 liter bottle (5.3.23)

- Add 500 ml of Milli-Q water (5.2.17)

- Weigh 7.3 grams of Sodium Chloride (5.2.7) and add to bottle.

- Weigh 2.383 grams of HEPES (5.2.9)

- Weigh 0.373 grams of Potassium Chloride (5.2.10)

- Weigh 0.144 grams of Calcium Chloride (5.2.11)

- Weigh 0.244 grams of Magnesium Chloride hexahydrate (5.2.12)

- Mix to dissolve all the compounds

- Adjust the pH to 7.4

- Fill up the bottle to the 1 Liter mark

- Store at 4C. (stable for 3 month)

### Stock solutions of test compounds

- Prepare stock solutions of the test compounds in a concentration of 10 mM in desired solvent. (Water, Methanol or DMSO)

### ATP quantification kit

- CellTiter-Glo® luminescent cell viability assay kit (Promega, G7572; 100 mL).

### Complete uptake buffer with 5.6 mM D-glucose

- On day 3 (i.e. day of exposure) of the testing procedure to be freshly made

- Unfreeze 1 Eppendorf tube with the 1 M D-Glucose solution (5.4.8)

- Take a 50 ml Falcon tube (5.3.19)

- Add 50 ml uptake buffer (5.4.9)

- Add 280 µL of 1 M D-Glucose (5.4.8)

- Mix the tube (use on day when produced)

### 10x concentrated test solutions in uptake buffer

- On day 3 (i.e. day of exposure) of the testing procedure to be freshly made

- Prepare in complete uptake buffer (5.4.12) a 10-fold stock solution in Eppendorf tubes (5.3.17) of the test chemicals, solvent control and reference chemical indocyanine green (ICG) (5.4.10) (f.c. 100 µM; 1% solvent)

### I-T4 working solution

- On day 4 (i.e. day of exposure) of the testing procedure to be freshly made in laboratory fit for working with radioactive chemicals!!!

- First prepare a stock solution of 100 nM T4 (f.c. in the assay will be 10 nM) in uptake buffer (5.4.9) from the 10 mM T4 solution (5.4.4)

- Subsequently add an appropriate amount of 125I-T4 tracer (5.1.5) to the stock solution of 100 nM T4 (f.c. 10000 CPM)

- Always test how much CPM is the solution by measuring 10 µL on the gamma counter (5.3.1) and adjust accordingly if necessaryProcedure

# Procedure

## In short

CHO-K1 OATP1C1 cells (with passage number in the range 41-54) are cultured in complete F12 media supplemented with G418 (f.c. 100 µg/mL) at 37°C with 5% CO2 saturation until 80% confluency, i.e. subcultured two times per week. On day 1 (72 hours before starting the TH uptake assay), CHO-K1 OATP1C1 cells are seeded at a density of 3x104 cells in 100 µL medium per well in a 96 wells plate. On day 3 (24 hours before start of the TH uptake assay), medium from CHO-K1 OATP1C1 assay plates is changed to serum-free F12 medium containing 5mM sodium-butyrate to increase expression of transfected SV-40 promotor plasmids. On day 4, the actual uptake experiment starts with a simultaneous 30 min exposure of the cells to the test compound and a 100 nM T4 solution containing radiolabeled 125I-T4 tracer. After exposure, cells are washed and lysed, and radioactivity of the lysate is determined as a measure for T4 uptake. Cell viability is assessed in a parallel exposure plate through the quantitation of ATP present in the cells.

## Day 1: Plating the cells

- Work sterile and in a flow cabinet

- Pre warm DMEM:F12 media (5.4.1), 1xPBS (5.4.5) and trypsin solution (5.4.7) in water bath (5.3.5) at 37C

- Take a sufficient amount of culture media

- Take a culture flask (80% confluent) with CHO-K1 OATP1C1 cells from the CO2 incubator

- Remove the media from the culture flask

- Rinse the cell layer with 10 mL 1xPBS (5.4.5)

- Remove the PBS

- Repeat the rinsing step 2x

- Add 2 mL trypsin solution (5.4.6)

- Spread it over the cell layer by canting the flask

- Remove 1 mL trypsin from the culture flask

- Place culture flask in the CO2 incubator (5.3.6)

- Incubate the flask for 15-30 minutes until the cells come lose from the bottom of the flask

- Suspend the lose cells in 10 mL media by pipetting up and down with a serological pipet (5.3.18)

- Count the cells with a cell counter (5.3.7)

- Dilute the cells with prewarmed F12 complete media (5.4.1) to 3x105 cells per mL

(You need at least 6 mL of suspension per 96 wells plate)

- Take a transparent (5.3.13) and a white (5.3.14) tissue culture treated 96 wells plate (sterile) and add 100 l of the cell suspension in the inner 60 wells of each plate leaving the outer row empty.

- Add 200 l 1xPBS (5.4.5) in the outer row wells of the plates.

- The transparent plate will be used for the 125I-TH uptake assay and the white plate for the cell viability test

- Place the filled 96 wells plate in the CO2 incubator (5.3.6) and incubate for 48 hours.

## Day 3: Changing to serum-free medium

- Dilute the serum free F12 media supplemented with 50 mM sodium butyrate (5.4.2) 10x in serum free F12 media (5.2.2)

- Remove media from inner 60 wells of each 96 wells plate (6.2)

- To induce OATP1C1 transcription, add 100 l of this medium (f.c. 5mM sodium butyrate) to the 60 inner wells

- Put the 96 wells plate in the CO2 incubator (5.3.6) and incubate for 24 hours.

## Day 4: performing the TH uptake assay

- Preheat the 1xPBS (5.4.5) and uptake buffer (5.4.12) at 37°C

- Take a transparent 96 wells plate with CHO-K1 OATP1C1 cells (6.3) from cell culture incubator

- Remove medium from the wells

- Wash once with 200 µL warm (37°C) 1xPBS (5.4.5)

- Add 80 µL of uptake buffer to each well (5.4.10)

- All further steps under 6.4 should be performed in a laboratory fit for working with radioactive chemicals!!!

- Add to the wells either 10 µL of test chemical (TC), solvent control (SC, f.c. 0.1%), or 10 µL reference compound SY dissolved in uptake buffer (5.4.13), or add uptake buffer itself (5.4.12) (blank) according to the plate layout in Figure 1

- Add 10 µL of 125I-T4 solution (5.4.14) to each well

- All waste after this step should be treated as radioactive waste!

|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A | X | X | X | X | X | X | X | X | X | X | X | X |
| B | X | SC TC1 | 0.03 µM TC1 | 0.1 µM TC1 | 0.3 µM TC1 | 1 µM TC1 | 3 µM TC1 | 10 µM TC1 | 30 µM TC1 | 100 µM TC1 | Blank | X |
| C | X | SC TC1 | 0.03 µM TC1 | 0.1 µM TC1 | 0.3 µM TC1 | 1 µM TC1 | 3 µM TC1 | 10 µM TC1 | 30 µM TC1 | 100 µM TC1 | Blank | X |
| D | X | SC TC1 | 0.03 µM TC1 | 0.1 µM TC1 | 0.3 µM TC1 | 1 µM TC1 | 3 µM TC1 | 10 µM TC1 | 30 µM TC1 | 100 µM TC1 | Blank | X |
| E | X | SC TC2 | 0.03 µM TC2 | 0.1 µM TC2 | 0.3 µM TC2 | 1 µM TC2 | 3 µM TC2 | 10 µM TC2 | 30 µM TC2 | 100 µM TC2 | 100 µM ICG | X |
| F | X | SC TC2 | 0.03 µM TC2 | 0.1 µM TC2 | 0.3 µM TC2 | 1 µM TC2 | 3 µM TC2 | 10 µM TC2 | 30 µM TC2 | 100 µM TC2 | 100 µM ICG | X |
| G | X | SC TC2 | 0.03 µM TC2 | 0.1 µM TC2 | 0.3 µM TC2 | 1 µM TC2 | 3 µM TC2 | 10 µM TC2 | 30 µM TC2 | 100 µM TC2 | 100 µM ICG | X |
| H | X | X | X | X | X | X | X | X | X | X | X | X |

Figure 1: Plate layout for the 125I-T4 uptake experiment. TC1 is test compound 1, TC2 is test compound 2. SC is solvent control (0.1%), blank is uptake buffer only, ICG is indocyanine green, a reference OATP1C1 inhibitor

- Incubate the plate at 37°C for 30 minutes.

- After incubation remove the incubation solution

- Wash each well once times with 150 µL ice-cold PBS + 0.1%BSA (5.4.6)

- Add 50 µL 0.1M NaOH (5.4.2) to each well

- Shake the plate at RT for 20 minutes (~700 RPM) (5.3.11)

- Transfer cell lysate from each well together with the pipet tip used for this transfer into separate 3.5 mL polypropylene tubes (5.3.21) , and close the tube with a push cap (5.3.22)

- Also transfer 10 µL of the 125I-T4 tracer solution (5.4.14) into a separate tube. This will be the 100% reference value.

- Mark the tubes and transfer the tubes in a rack for the gamma counter (5.3.1). A holder can hold exactly 10 tubes, so you can measure the 60 wells of the plate per row!

- Measure radioactivity on the gamma counter (5.3.1).

## Day 4: performing the cell viability test

- Preheat the 1xPBS (5.4.5) and uptake buffer (5.4.12) at 37°C

- Take a transparent 96 wells plate with CHO-K1 OATP1C1 cells (6.3) from cell culture incubator

- Remove medium from the wells

- Wash once with 200 µL warm (37°C) 1xPBS (5.4.5)

- Add 80 µL of uptake buffer to each well (5.4.12)

- Add to the wells either 10 µL of test chemical (TC), solvent control (SC, f.c. 0.1%), or 10 µL reference compound SY dissolved in uptake buffer (5.4.13), or add uptake buffer itself (5.4.11) (blank) according to the plate layout in Figure 1

- Add 10 µL of T4 solution (5.4.4) to each well.

- Incubate the plate at 37°C for 30 minutes.

- After incubation remove the incubation solution

- Wash each well once with 150 µL cold PBS + 0.1%BSA (5.4.6)

- Add 100 μL of the CellTiter-Glo® Reagent (5.4.10) to each well

- Shake the plates on an orbital shaker for 2 min (~450 RPM) (5.3.11)

- Incubate at room temperature for 10 min

- Measure luminescence in a microplate reader (5.3.2)

# Data analysis

## TH uptake assay

- Import the radioactivity readouts from the gamma counter into Microsoft Excel

- Correct each readout by dividing it with the value from the 125I-T4 tracer solution (100% reference value)

- Normalise the average value of each treatment concentration against the median value of the solvent control responses (n=6 if both test compounds are dissolved in the same solvent, else n=3 (see Figure 1)) and against the maximum inhibition observed for the reference compound ICG (n=3), according to the formula

Yi = (RAi – RAICG_median)/(RASC_median – RAICG_median),

with Yi is the normalized T4 uptake in measurement i, RAi is the radioactivity for measurement i, RAICG_median is the median radioactivity measured for the reference chemical ICG (minimum uptake), and RASC_median is the median radioactivity of the negative solvent control (maximum uptake).

- Multiply the normalized values by 100 in order to obtain the % T4 uptake inhibition in each treatment concentration.

- In case the three lowest treatment concentrations provide evidence for a constant plateau response which is significantly different from the 100% control reference, re-normalise the normalised effect data to the average response of the three lowest test compound concentrations  (doi:10.14573/altex.1803231; doi.org/10.14573/altex.2210171)

- Transfer the normalized effect data (% of T4 uptake in each concentration) to an appropriate statistical data analysis software for dose-response curve fitting and the estimation of a benchmark concentration (BMC) or inhibitory concentration (IC), e.g. GraphPad Prism (http://www.graphpad.com), or alternatively, the BMC calculation script in the R language (https://doi.org/10.1371/journal.pone.0146021, https://cran.r-project.org/web/packages/drc/index.html).

- Fit a concentration-response curve by performing a non-linear regression on the normalized cellular uptake of T4 (Y) as a function of the exposure concentration (C) according to the sigmoidal 4-parametric Hill model

Y=Bottom+Top-Bottom1+(IC50/C)r

with Top and Bottom being the maximum and minimum asymptotes of the sigmoidal concentration-response curve (in percentages), respectively, IC50 being the exposure concentration that gives a T4 uptake half way between Bottom and Top (which is not necessarily the same as the response at Effect=50%), and r being the Hill slope describing the steepness of the curve.

- Determine the benchmark concentration BMC20 corresponding to 20% inhibition of T4 uptake (i.e. 80% remaining uptake compared to the negative control) as

BMC20=IC50rA    , with A=Top-Bottom80-Bottom-1.

The 20% inhibition level is proposed as optimal low effect level for a statistically robust BMC estimation because it corresponds to the 1.5 fold standard deviation of the inter-study variability (derived from more than 20 chemical exposure studies with each of three independent repeats).

## Cell viability assay

- Import the fluorescence readouts from the platereader into Microsoft Excel

- Normalise the average value of each treatment concentration against the median value of the solvent control responses (n=6 if both test compounds are dissolved in the same solvent, else n=3 (see Figure 1)):

Zi = (Fi)/(FSC_median)

with Zi is normalized cell viability in measurement i, Fi is the fluorescence of measurement i and FSC_median is the median fluorescence of the negative solvent control (no cytotoxicity).

- Multiply the normalized values by 100 in order to obtain the % cell viability in each treatment concentration.

- Transfer the normalized cell viability results to an appropriate statistical data analysis software for dose-response curve fitting and the estimation of a benchmark concentration (BMC) or inhibitory concentration (IC), e.g. GraphPad Prism (http://www.graphpad.com), or alternatively, the BMC calculation script in the R language (https://doi.org/10.1371/journal.pone.0146021, https://cran.r-project.org/web/packages/drc/index.html).

- Fit a concentration-response curve by performing a non-linear regression on the normalized cell viability (Z) as a function of the exposure concentration (C) according to a sigmoidal 3-parametric Hill model with constrained minimum of zero (Bottom = 0)

Z=Top1+(IC50/C)r

with Top and zero being the maximum and minimum asymptotes of the sigmoidal concentration-response curve (in percentages), respectively, IC50 being the exposure concentration that gives a cell viability response half way between Bottom and Top (which is not the same as the response at Effect=50%), and r being the Hill slope describing the steepness of the curve.

- Determine the benchmark concentration BMC10 corresponding to 10% inhibition of cell viability (i.e. 90% remaining uptake compared to the negative control) as

BMC10=IC50rB    , with B=Top90-1.

- Decreases in T3 uptake should not be considered as specific inhibition of OATP1C1 if cell viability is decreased by more than 10% (i.e. at concentrations exceeding the BMC10 for cell viability.

## Acceptance criteria

### The T4 uptake in all three replicate negative solvent controls should be 10-30% of the radioactivity of the 125I-T4 tracer solution added to each well.

### The T4 uptake in all three replicate exposures to the positive reference compound ICG should be <8% of the radioactivity of the 125I-T4 tracer solution added to each well.

### The normalized T4 uptake in all three replicate exposures to reference compound ICG should

### be <40% of the normalized T4 uptake after exposure to the negative solvent control.

### BMC20 for T4 uptake can only be reported as specific for OATP1C1 inhibition if it is below the BMC10 for cell viability.

# References

Pizzagalli F, Hagenbuch B, Stieger B, Klenk U, Folkers G, Meier PJ. 2002 Identification of a novel human organic anion transporTIng polypepTIde as a high affinity thyroxine transporter. Mol Endocrinol 16:2283–2296. https://doi.org/10.1210/me.2001-0309

Wagenaars F, Cenijn P, Chen Z, Meima M, Scholze M, Hamers T. 2024. Two novel in vitro assays to screen chemicals for their capacity to inhibit thyroid hormone transmembrane transporter proteins OATP1C1 and OAT4. Archives of Toxicology 98: 3019-3034. https://doi.org/10.1007/s00204-024- 03787-2